## Velocity tests
Velocity tests
Date: |  | 2023-11-03 |  | Theoretical distance travelled (mm): |  |  |  | 72
Room temperature: |  | 20.5°C
Humidity: |  | 48.7%RH
Trial | Central ball spline [Fixed / Mobile] | Angular velocity (rad/s) | mass (g) | mass mean (g) | std mass (g) | Total measured traveled distance (mm) | Slip ratio | Slip ratio mean | slip ratio std | Measred traveled distance mean (mm) | Measured traveled distance std (mm) | Slip ratio mean per ball spline condition
BF-V1-15%-1 | Fixed | 1π | 99 | 99.8 | 0.74833 | 29 | 0.59722 | 0.54167 | 0.051967 | 33 | 3.7417 | 0.61265
BF-V1-15%-2 | Fixed | 1π | 100.8 |  |  | 38 | 0.47222
BF-V1-15%-3 | Fixed | 1π | 99.6 |  |  | 32 | 0.55556
BF-V2-15%-1 | Fixed | 2π | 100.1 | 100.0333 | 0.24944 | 21 | 0.70833 | 0.63889 | 0.06314 | 26 | 4.5461
BF-V2-15%-2 | Fixed | 2π | 99.7 |  |  | 32 | 0.55556
BF-V2-15%-3 | Fixed | 2π | 100.3 |  |  | 25 | 0.65278
BF-V3-15%-1 | Fixed | 3π | 99.1 | 99.7 | 0.77889 | 26 | 0.63889 | 0.65741 | 0.017322 | 24.6667 | 1.2472
BF-V3-15%-2 | Fixed | 3π | 100.8 |  |  | 25 | 0.65278
BF-V3-15%-3 | Fixed | 3π | 99.2 |  |  | 23 | 0.68056
BM-V1-15%-1 | Mobile | 1π | 100.2 | 99.6333 | 0.44969 | 42 | 0.41667 | 0.35185 | 0.047213 | 46.6667 | 3.3993 | 0.31636
BM-V1-15%-2 | Mobile | 1π | 99.1 |  |  | 50 | 0.30556
BM-V1-15%-3 | Mobile | 1π | 99.6 |  |  | 48 | 0.33333
BM-V2-15%-1 | Mobile | 2π | 99 | 99.5 | 0.57155 | 48 | 0.33333 | 0.30556 | 0.060007 | 50 | 4.3205
BM-V2-15%-2 | Mobile | 2π | 99.2 |  |  | 46 | 0.36111
BM-V2-15%-3 | Mobile | 2π | 100.3 |  |  | 56 | 0.22222
BM-V3-15%-1 | Mobile | 3π | 100.3 | 99.5667 | 0.52493 | 52 | 0.27778 | 0.29167 | 0.051967 | 51 | 3.7417
BM-V3-15%-2 | Mobile | 3π | 99.3 |  |  | 46 | 0.36111
BM-V3-15%-3 | Mobile | 3π | 99.1 |  |  | 55 | 0.23611
Set up
The central ball spline is fixed with 4 mm cable clips. So it can only rotate in place. Gelatin was made the day before by mixing 300 g gelatin with 1700 mL tap water (Rotterdam, Zuid-Holland). Gelatin is cut into pieces of 40x100 mm. Gelatin is weighted.
Procedure
1. needle is inserted 40 mm, with the hand as close by the gelatine as possible to get a straight line
2. needle and cart are aligned
3. metronome start
4. Rotate crank 15 cycles
π 2π 3π

## Gelatin concentration tests
Gelatin concentration tests
Date: | 2023-11-01 |  | Theoretical distance travelled (mm): |  |  |  | 72
Room temperature: | 21°C
Humidity: | 53.4%RH
Trial | Gelatin concetration (wt%) | mass (g) | mass mean (g) | std mass (g) | Total measured traveled distance (mm) | Slip ratio | Slip ratio mean | slip ratio std | Measred traveled distance mean (mm) | Measured traveled distance std (mm)
BM-V1-5%-1 | 5 | 94 |  |  | 59 | 0.18056 | 0.18056 | 0.01134 | 59 | 0.8165
BM-V1-5%-2 | 5 | 94 |  |  | 58 | 0.19444
BM-V1-5%-3 | 5 | 94 |  |  | 60 | 0.16667
BM-V1-10%-1 | 10 | 72.2 |  |  | 58 | 0.19444 | 0.18981 | 0.0065473 | 58.3333 | 0.4714
BM-V1-10%-2 | 10 | 72.2 |  |  | 59 | 0.18056
BM-V1-10%-3 | 10 | 72.2 |  |  | 58 | 0.19444
BM-V1-10%-1 | 15 | 100.2 |  |  | 42 | 0.41667 | 0.35185 | 0.047213 | 46.6667 | 3.3993
BM-V1-10%-2 | 15 | 99.1 |  |  | 50 | 0.30556
BM-V1-10%-3 | 15 | 99.6 |  |  | 48 | 0.33333
5 10 15

## Needle insertion force tests
Needle insertion force tests
Date: |  | 2023-10-27
Room temperature: |  | 22.1 °C
Humidity: |  | 51.3 %RH
Trial | Gelatine (wt%) | Insertion force (N) | Force mean (N) | Force std (N)
A-5%-1 | 5 | 0.25 | 0.15 | 0.08165
A-5%-2 | 5 | 0.15
A-5%-3 | 5 | 0.05
A-10%-1 | 10 | 0.75 | 0.8 | 0.1472
A-10%-2 | 10 | 0.65
A-10%-3 | 10 | 1
A-15%-1 | 15 | -
A-15%-2 | 15 | -
A-15%-3 | 15 | -
Set up
Lasercutting box fixed to Thorlabs breadboard with 4 bolts. Spirit level. Aluminum festo beam (length 282 mm, height/width 30mm) to support the needle and prevent buckling of the unsupported length. PMMA plate of 2mm under the tissue box to get to the same height. Push and pull force gauge tie wraps to 2 cross shaped plates that can slide into the festo beam (festo mini slide) (length 180 mm, height/width 30mm)
Procedure
1. Needle is inserted 20 mm.
2. Peak force button pushed to set at zero
3. Calibrate force gauge
4. Force gauge is pushed. Then pushed for another 80mm by pushing the gauge meter forward by hand over the slider.
5. Peak force is read from display and noted
3 times, each time a new trail in the gelatin
Observations/Notes
Insertion in 15 wt% was not possible because of too much bending of the needle

## Actuation system force tests
Actuation system force tests
Date: |  | 2023-11-02
Room temperature: |  | 21.0 °C
Humidity: |  | 44.5 %RH
Trial | Force (N) | Force mean (N) | Force std (N)
1 | 0.15 | 0.175 | 0.12042
2 | 0.1
3 | 0.05
4 | 0.4
5 | 0.175
Set up
The gauge meter is heightend with PMMA plates 2 of 6mm and 2 of 2mm a total of 16 mm. The plates and gauge are taped together. Two bolts are screwed behind the PMMA tower to prevent the gauge meter from sliding.
Procedure
1. Prototype is moved forward so 1 needle segment touches the force gauge.
2. The crank is turned counter clockwise
3. The max force is measured

## Theoretical critical load
Theoretical critical load computation
Design parameter |  | unit
E_steel | 200000 | N mm^-2
r_needlesegment | 0.15 | mm
I | 0.00039761 | mm^4
K | 2
L_substrate | 4 | mm
mu | 0.008 | N mm^-2
L_air | 20 | mm
F_cr ext | 0.5035 | N